MiSeq V3 Kit Sample Preparation, August 2014

Ideal sample libraries will be standardized to 10nM. Based on the KAPA quantification in the previous step, determine if libraries need to be diluted or concentrated for a working 10nM concentration.

*Before beginning, set a heat block to 96ºC for last-step heat shock
*If necessary, complete the wash steps on the MiSeq to prepare for sequencing
*Remove HT1 from kit and let thaw, but keep chilled. Remove loading tray from -20 freezer and let thaw in a bucket of tepid water in a dark room. (Don't do this if another qPCR is required)

Dilution
1. Determine appropriate volumes of EBT (EB buffer with Tween) to add to a 10μL aliquot of prepared libraries from qPCR results. Aliquot 10μL of library as precisely as possible into a separate tube. Pipet past the stop.

2. Add volumes of EBT to each tube as precisely as possible. Vortex and spin down 2X.

Concentration
1. Precisely determine volume left in tube via pipetting. This will determine the final volume of the reconstituted sample.

2. Use vacuum centrifuge to dry samples using the V-AQ setting. For small volumes of DNA 30ºC is sufficient; drying times are dependent on the volume.

3. Add appropriate amounts of EBT to each tube. Because the volume will be small (<20μL), vortex carefully to keep liquid in bottom of tube. Spin down and repeat vortexing.

Pooling
Pool all normalized libraries together to make a working stock of 10nM pooled. The volume of the smallest sample will determine the volume to add (e.g. if the smallest sample is 5μL, add 5μL of each library to the pool). Optional: Redo KAPA library quantification qPCR step.

Library Denaturation
1. Combine the following volumes of DNA and freshly diluted 0.1 N NaOH in a 1.5μL tube:
- 5 μL 10 nM sample DNA
- 5 μL 0.1 N NaOH
Vortex briefly to mix and spin down.

2. Incubate library for 5 minutes at room temperature to denature DNA.

3. Add 990 μL of pre-chilled HT1 to the tube. This results in a 50pM denatured library in 1 mM NaOH.

Library Dilution
Different library concentrations are required for optimal cluster generation. These can range anywhere from 12 to 50 pM, depending on the chemistry. The following table shows the appropriate volumes of the 50pM denatured library and pre-chilled HT1 to combine for 600μL of the desired concentrations.

Final library concentration (pM)	15pM	18pM	20pM	25pM
Volume of 50 pM denatured DNA	180	216	240	300
Volume of pre-chilled HT1	420	384	360	300

Vortex the mixture and spin down. Place on ice until ready for the final heat shock and loading. The final input volume is 600μL. If a control is desired (PhiX) it must be in the same final concentration as the DNA and must be added within the final volume of 600μL for loading.

Preparing the machine
1. Ensure all washes have completed. If the machine has not finished a required wash, it will not let you begin a run.

2. Complete the MiSeq sample sheet using the Illumina Experiment Manager (in a Windows machine only) and save to an external to transport it to the machine. Confirm sample sheet is correct and save it to the machine under the specified run name.

3. Follow instructions on machine for run setup. Perform final heat shock of libraries by putting in 96ºC heat block (taped shut) for 2 minutes. Remove immediately and place in ice/water bath to chill for at least 5 minutes. Keep on ice until completely ready to load.

4. Check reagent tray for complete thawing (no ice crystals hitting sides). Remove flow cell from storage container and rinse THOROUGHLY with diH20. Dry with Kimwipe and wipe down with alcohol pad, finishing off with a polish with lens paper. Insert flow cell, putting old cell into storage container. Pierce foil covering input well on reagent tray with a P1000 tip and load library.

5. Ensure all components are assembled and that the machine begins the run. BaseSpace will monitor the entire run until completion.